## Slide 1
Through Vial Impedance Spectroscopy (TVIS)
Dual-electrode System for Process Parameter Determination
Geoff Smith
Leicester School of Pharmacy, De Montfort University, United Kingdom
ISL-FD Midwest Chapter Annual Meeting, April 12th 2018
Midwest Conference Centre, Northlake, IL 60164, USA
www.dmu.ac.uk/tvis

## Slide 2: Outline
Description of TVIS measurement system
Applications in Brief
First time report on the use of dual-electrode system and its applications
    Ice region specific temperature prediction (, )
    Drying rate determination
    Heat transfer coefficient () determination
Acknowledgements
TVIS dielectric loss mechanisms
Notes: Change picture for you at freeze-dryer

## Slide 3
Through Vial Impedance Spectroscopy (TVIS)
Description of Measurement System

## Slide 4: Introduction to the TVIS System
Impedance spectroscopy characterizes the ability of materials to conduct electricity under an applied an oscillating voltage (of varying frequency)
Impedance measurements across a vial rather than within the vial
Hence “Through Vial Impedance Spectroscopy”
Features
  Single vial “non-product invasive”
  Both freezing and drying characterised in a single technique
  Non-perturbing to the packing of vials
  Stopper mechanism unaffected
TVIS

## Slide 5
Freeze drying chamber
Stimulating voltage
Resultant current
LyoDEATM measurement software
Junction
box
TVIS system
(I to V convertor)
Pass-through
TVIS
measurement vial
LyoViewTM analysis software
Notes: Version 14 change : don’t animate first picture – make it appear with the slide
Change photo for the pass-through to the new one
V4 edited slide

## Slide 6
Through Vial Impedance Spectroscopy (TVIS)
Applications

## Slide 7: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 8: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 9: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
increase
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Surrogate drying rate (from )
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 10: Through Vial Impedance Spectroscopy (TVIS)
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034
Time 0.3 h | 0.27396 | 0.24912 | 0.22663 | 0.2063 | 0.18797 | 0.17147 | 0.15668 | 0.14349 | 0.13183 | 0.12162 | 0.11284 | 0.10548 | 0.099574 | 0.095186 | 0.092431 | 0.091474 | 0.092551 | 0.095973 | 0.10215 | 0.11162 | 0.12501 | 0.14319 | 0.16715 | 0.198 | 0.23699 | 0.28544 | 0.34395 | 0.41214 | 0.48865 | 0.56683 | 0.63782 | 0.69048 | 0.71191 | 0.69725 | 0.65001 | 0.57934 | 0.49898 | 0.41726 | 0.34526 | 0.28116 | 0.22886 | 0.17991 | 0.14643 | 0.1181 | 0.089415 | 0.072054
Time 4 h | 0.099081 | 0.096153 | 0.095569 | 0.097591 | 0.10258 | 0.11101 | 0.12349 | 0.14075 | 0.16366 | 0.19315 | 0.23007 | 0.27492 | 0.32735 | 0.38533 | 0.44432 | 0.49685 | 0.53348 | 0.54588 | 0.53067 | 0.49125 | 0.43601 | 0.37424 | 0.31336 | 0.25794 | 0.20992 | 0.16948 | 0.13623 | 0.10926 | 0.087318 | 0.069838 | 0.055881 | 0.044566 | 0.035658 | 0.0285 | 0.02283 | 0.018246 | 0.014616 | 0.011669 | 0.0094101 | 0.0075633 | 0.0061328 | 0.0048349
Time 0.3 h | 0.26705 | 0.24478 | 0.21978 | 0.1979 | 0.17863 | 0.15615 | 0.14276 | 0.12639 | 0.11756 | 0.11024 | 0.10127 | 0.1004 | 0.098391 | 0.10089 | 0.10904 | 0.11009 | 0.11451 | 0.12496 | 0.13439 | 0.14683 | 0.16234 | 0.18119 | 0.20238 | 0.22944 | 0.26256 | 0.30331 | 0.35113 | 0.4089 | 0.47307 | 0.54005 | 0.60366 | 0.65431 | 0.67974 | 0.67269 | 0.63391 | 0.56955 | 0.49367 | 0.41413 | 0.34323 | 0.27998 | 0.22777
Time 4 h | 0.093429 | 0.092843 | 0.096342 | 0.10075 | 0.11232 | 0.12322 | 0.1435 | 0.1596 | 0.19684 | 0.23027 | 0.27214 | 0.31647 | 0.36246 | 0.41025 | 0.45421 | 0.48896 | 0.51147 | 0.51695 | 0.50285 | 0.46926 | 0.42098 | 0.36431 | 0.30625 | 0.25218 | 0.20511 | 0.16565 | 0.13355 | 0.10767 | 0.086605 | 0.069689 | 0.056026 | 0.045028 | 0.036377 | 0.029122 | 0.023205 | 0.018299 | 0.014688 | 0.011778 | 0.0095536 | 0.0077519 | 0.0063112
FPEAK
FPEAK
Liquid State
Solid State
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
increase
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
Monitoring Phase Behaviour (ice nucleation temperature and solidification end points by using
temperature calibration for predicting temperature of the product in primary drying
Surrogate drying rate (from )
  is highly sensitive to low ice volumes; therefore it could be used for determination end point of primary drying
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 11: TVIS Response Surface (3D-Plot)
Imaginary Part of Capacitance
Real Part of Capacitance
High frequency
Liquid state
Liquid state
Frozen solid
Frozen solid
low frequency
Annealing = Re-heating and Re-cooling
Re-heating
Re-heating
Intermediate frequency
Re-cooling
Re-cooling
low frequency
Primary drying
Primary drying
low frequency
Notes: Full Shelf Water Study_010716_Ch1 (Filling volume = 2 mL) touch base electrode

## Slide 12: Dielectric loss spectrum
Data analysing software (LyoView ™) identifies the peak frequency () and peak amplitude () in the imaginary part of the capacitance spectrum
Chart: Log Frequency
Chart: Log Frequency

## Slide 13
Through Vial Impedance Spectroscopy (TVIS)
Dual-electrode system and its applications
(Ice temperature, Drying rate and Heat transfer coefficient )

## Slide 14: Dual-electrode system
Electrode height ~10 mm
Electrode distance from base ~3 mm
Standard TVIS vial
(Single electrode system)
Electrode Dimension
10 x 19 mm
New feature of TVIS vial
(Dual electrode system)
Electrode Dimension
Top electrode (TE): 10 x 19 mm
Bottom Electrode (BE): 5 x 19 mm
Bottom electrode height ~10 mm
Electrode distance from base ~3 mm
Top electrode height ~15 mm
Gap between electrode ~3 mm
A dual electrode system comprises two pairs of copper electrode glued to the external surface of a Type I tubular glass vial.
This option is suitable for large volume samples, including those used for determination.
~3.4 g of water
~8 g of water

## Slide 15: Temperature Determination
: TVIS predicted temperature
from top electrode (TE)
: TVIS predicted temperature
from bottom electrode (BE)
Both and can be estimated by extrapolating from the temperatures predicted from the centers of top electrode () and bottom electrode ().
Dual electrode
3 mm gap
3 mm gap
5 x 19 mm
15 x 19 mm

## Slide 16: Aims & Objectives
Aims
To determine the heat transfer coefficient () by using a novel dual electrode TVIS approach
Temperature calibration of log of top electrode () and bottom electrode ()
I
Prediction ice temperatures for both electrodes during primary drying
II
Temperature calibration of
III
Compensation of during primary drying
IV
Calibration of for ice layer height
V
Estimation of ice layer height during primary drying
VI
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period
IX
Heat transfer coefficient ( calculation
X

## Slide 17: Objective
Temperature calibration of log of top electrode () and bottom electrode ()
I
Annealing the sample
In-line TVIS measurement
Identifying peak frequency () using LyoView ™ software
Identifying peak frequency (𝐹_𝑃𝐸𝐴𝐾) using LyoView ™ software
Predicting product temperature using calibration plot
Calibration plot (temperature vs Log )
Calibration plot (temperature vs Log 𝐹_𝑃𝐸𝐴𝐾)
Annealing the sample
In-line TVIS measurement
Predicting product temperature using calibration plot

## Slide 18: Objective
Chart: Time /h
Shelf Temp.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 6.5 h | -0.011026 | -0.012215 | -0.027689 | -0.023597 | -0.0069787 | -0.0053348 | -0.0030174 | 0.0013325 | 0.0065751 | 0.0097044 | 0.013438 | 0.017118 | 0.021618 | 0.026952 | 0.033746 | 0.041929 | 0.052016 | 0.064186 | 0.079127 | 0.097201 | 0.11944 | 0.146 | 0.1783 | 0.2178 | 0.26517 | 0.32227 | 0.39079 | 0.47092 | 0.56228 | 0.66169 | 0.76175 | 0.85177 | 0.91802 | 0.94907 | 0.93932 | 0.8934 | 0.82052 | 0.72661 | 0.63632 | 0.54819 | 0.46021 | 0.40033 | 0.33842 | 0.3044 | 0.2558 | 0.23181 | 0.21255 | 0.20425 | 0.20259 | 0.20868 | 0.22018
Time 6.6 h | -0.010844 | -0.013094 | -0.028215 | -0.023415 | -0.0063289 | -0.0045493 | -0.0018557 | 0.0026925 | 0.0080769 | 0.011546 | 0.015785 | 0.020036 | 0.025265 | 0.031534 | 0.039471 | 0.049058 | 0.06099 | 0.075285 | 0.092893 | 0.11431 | 0.14093 | 0.1727 | 0.21171 | 0.25948 | 0.31676 | 0.38554 | 0.46659 | 0.55908 | 0.65975 | 0.76252 | 0.85551 | 0.92559 | 0.96052 | 0.95286 | 0.90669 | 0.83273 | 0.74011 | 0.65262 | 0.55298 | 0.4715 | 0.40258 | 0.34125 | 0.29102 | 0.26019 | 0.22652 | 0.21129 | 0.1945 | 0.20035 | 0.20414 | 0.21607 | 0.23162
Time 6.7 h | -0.010233 | -0.013175 | -0.028274 | -0.022743 | -0.0054069 | -0.0034808 | -0.00042223 | 0.0044144 | 0.01012 | 0.014066 | 0.018998 | 0.024035 | 0.030298 | 0.037789 | 0.047292 | 0.058826 | 0.073223 | 0.090722 | 0.11222 | 0.13839 | 0.17098 | 0.21029 | 0.25838 | 0.3172 | 0.3871 | 0.46963 | 0.56416 | 0.66707 | 0.7712 | 0.86582 | 0.93655 | 0.97094 | 0.96234 | 0.91434 | 0.83635 | 0.74343 | 0.64608 | 0.55833 | 0.47266 | 0.39948 | 0.33451 | 0.29245 | 0.25353 | 0.23378 | 0.20871 | 0.19984 | 0.1997 | 0.2015 | 0.21198 | 0.22733 | 0.24895
Time 6.8 h | -0.0099355 | -0.010961 | -0.025677 | -0.02165 | -0.0043092 | -0.0020771 | 0.001197 | 0.0065657 | 0.012802 | 0.017476 | 0.023499 | 0.029632 | 0.037339 | 0.046572 | 0.058415 | 0.072736 | 0.09063 | 0.11238 | 0.13935 | 0.17222 | 0.21322 | 0.26261 | 0.32283 | 0.39563 | 0.48042 | 0.57735 | 0.68277 | 0.7887 | 0.88235 | 0.9508 | 0.98058 | 0.96664 | 0.91424 | 0.83246 | 0.73667 | 0.63856 | 0.5471 | 0.46601 | 0.38946 | 0.33579 | 0.28239 | 0.25179 | 0.22417 | 0.21528 | 0.19822 | 0.19607 | 0.20091 | 0.21064 | 0.22634 | 0.24719 | 0.27274
Time 6.9 h | -0.0089346 | -0.010265 | -0.024797 | -0.020617 | -0.0027907 | -0.00025292 | 0.0035689 | 0.0097638 | 0.016718 | 0.022291 | 0.029711 | 0.037389 | 0.047034 | 0.05871 | 0.073704 | 0.091917 | 0.11467 | 0.14237 | 0.17678 | 0.21881 | 0.27107 | 0.33354 | 0.40861 | 0.49749 | 0.59727 | 0.70497 | 0.81179 | 0.90383 | 0.96632 | 0.98864 | 0.96657 | 0.90704 | 0.82145 | 0.72363 | 0.62477 | 0.53332 | 0.45336 | 0.38378 | 0.32751 | 0.28107 | 0.24243 | 0.21966 | 0.20661 | 0.19414 | 0.1959 | 0.20098 | 0.21192 | 0.22772 | 0.24915 | 0.2732 | 0.30165
Time 7 h | -0.0079524 | -0.0096422 | -0.024013 | -0.019274 | -0.00065527 | 0.0023502 | 0.0069457 | 0.014294 | 0.022332 | 0.029183 | 0.0386 | 0.048527 | 0.061028 | 0.076199 | 0.095709 | 0.11952 | 0.14917 | 0.18538 | 0.23021 | 0.28467 | 0.35158 | 0.43028 | 0.52255 | 0.62693 | 0.73646 | 0.8422 | 0.92939 | 0.98297 | 0.99229 | 0.95868 | 0.8898 | 0.79968 | 0.70081 | 0.60339 | 0.51292 | 0.43557 | 0.37221 | 0.3158 | 0.27694 | 0.24492 | 0.21667 | 0.20627 | 0.19841 | 0.20582 | 0.20492 | 0.21652 | 0.23273 | 0.2534 | 0.27813 | 0.30477 | 0.33401
Time 7.1 h | -0.0072225 | -0.0078733 | -0.021826 | -0.017104 | 0.0022508 | 0.0058654 | 0.011358 | 0.020219 | 0.029818 | 0.038452 | 0.050584 | 0.063481 | 0.079875 | 0.09973 | 0.12537 | 0.15659 | 0.19542 | 0.24264 | 0.30066 | 0.37027 | 0.45418 | 0.54985 | 0.65682 | 0.7684 | 0.8716 | 0.95211 | 0.9947 | 0.99196 | 0.94565 | 0.86934 | 0.77503 | 0.67549 | 0.5806 | 0.49486 | 0.41965 | 0.35895 | 0.30957 | 0.26685 | 0.23671 | 0.2202 | 0.21086 | 0.19933 | 0.20391 | 0.21158 | 0.22344 | 0.23974 | 0.26069 | 0.28443 | 0.31084 | 0.33736 | 0.36523
Time 7.2 h | -0.0058046 | -0.0061413 | -0.019458 | -0.014478 | 0.0061524 | 0.010622 | 0.01737 | 0.028317 | 0.039983 | 0.051176 | 0.06699 | 0.083978 | 0.10566 | 0.13187 | 0.16569 | 0.20672 | 0.25752 | 0.3187 | 0.39276 | 0.47969 | 0.58074 | 0.68993 | 0.80068 | 0.9005 | 0.97199 | 1.0024 | 0.98623 | 0.92908 | 0.84405 | 0.74734 | 0.64904 | 0.55798 | 0.47566 | 0.40517 | 0.34692 | 0.30148 | 0.26468 | 0.23973 | 0.21784 | 0.21142 | 0.20723 | 0.20722 | 0.21754 | 0.22399 | 0.24796 | 0.26887 | 0.29214 | 0.31742 | 0.3425 | 0.36748 | 0.39208
-25 °C
Chart: Log FPEAK
| 2.7097859999999998 | 2.7409289999999999 | 2.7741730000000002 | 2.8071109999999999 | 2.8420100000000001 | 2.8772090000000001 | 2.9142049999999999 | 2.9513449999999999 | 2.9901089999999999 | 3.0288729999999999 | 3.0676369999999999 | 3.1087199999999999 | 3.1495929999999999 | 3.1926139999999998 | 3.2352349999999999 | 3.2774969999999999 | 3.3203369999999999 | 3.3627950000000002 | 3.4058160000000002 | 3.4484370000000002 | 3.4906990000000002 | 3.5335390000000002 | 3.57369 | 3.6122730000000001 | 3.6463709999999998 | 3.6786080000000001
Top electrode | -40.6 | -40 | -39.1 | -38.3 | -37.4 | -36.6 | -35.5 | -34.8 | -33.9 | -33.1 | -32 | -31 | -30.2 | -29 | -28 | -27 | -26 | -25 | -24.1 | -23.1 | -22.2 | -21.1 | -20.2 | -19.3 | -18.5 | -17.9
Bottom electrode | -40.1 | -39.4 | -38.6 | -37.8 | -36.8 | -35.9 | -35 | -34.1 | -33.2 | -32.3 | -31.2 | -30.3 | -29.3 | -28.4 | -27.3 | -26.2 | -25.3 | -24.2 | -23.4 | -22.3 | -21.3 | -20.4 | -19.4 | -18.5 | -17.8 | -17.2
Bottom electrode
Top electrode
Top electrode
TE | -1.02 | 30.1 | -114
BE | -5.43 x 10-1 | 26.7 | -109
TE | -1.02 | 30.1 | -114
BE | -5.43 x 10-1 | 26.7 | -109
Polynomial coefficient from calibration
Temperature calibration of log of top electrode () and bottom electrode ()
I

## Slide 19: Objective
Prediction ice temperatures for both electrodes during primary drying
II

## Slide 20: Objective
Chart: Time / h
Shelf Temperature
Chart: Time / h
Temp. constant
Temp. constant
Prediction ice temperatures for both electrodes during primary drying
II
The product temperature predicted by TVIS can demonstrate the temperature gradient across ice cylinder height

## Slide 21: Objective
Temperature calibration of
III
Annealing the sample
In-line TVIS measurement
Identifying peak amplitude () using LyoView ™ software
Identifying peak amplitude (𝐶_𝑃𝐸𝐴𝐾^″) using LyoView ™ software
Temperature compensation of using calibration plot
Temperature compensation of 𝐶_𝑃𝐸𝐴𝐾^″ using calibration plot
Calibration plot ( vs temperature)
Calibration plot (𝐶_𝑃𝐸𝐴𝐾^″ vs temperature)
Annealing the sample
In-line TVIS measurement

## Slide 22: Objective
Chart: Time /h
Shelf Temp.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 6.5 h | -0.011026 | -0.012215 | -0.027689 | -0.023597 | -0.0069787 | -0.0053348 | -0.0030174 | 0.0013325 | 0.0065751 | 0.0097044 | 0.013438 | 0.017118 | 0.021618 | 0.026952 | 0.033746 | 0.041929 | 0.052016 | 0.064186 | 0.079127 | 0.097201 | 0.11944 | 0.146 | 0.1783 | 0.2178 | 0.26517 | 0.32227 | 0.39079 | 0.47092 | 0.56228 | 0.66169 | 0.76175 | 0.85177 | 0.91802 | 0.94907 | 0.93932 | 0.8934 | 0.82052 | 0.72661 | 0.63632 | 0.54819 | 0.46021 | 0.40033 | 0.33842 | 0.3044 | 0.2558 | 0.23181 | 0.21255 | 0.20425 | 0.20259 | 0.20868 | 0.22018
Time 6.6 h | -0.010844 | -0.013094 | -0.028215 | -0.023415 | -0.0063289 | -0.0045493 | -0.0018557 | 0.0026925 | 0.0080769 | 0.011546 | 0.015785 | 0.020036 | 0.025265 | 0.031534 | 0.039471 | 0.049058 | 0.06099 | 0.075285 | 0.092893 | 0.11431 | 0.14093 | 0.1727 | 0.21171 | 0.25948 | 0.31676 | 0.38554 | 0.46659 | 0.55908 | 0.65975 | 0.76252 | 0.85551 | 0.92559 | 0.96052 | 0.95286 | 0.90669 | 0.83273 | 0.74011 | 0.65262 | 0.55298 | 0.4715 | 0.40258 | 0.34125 | 0.29102 | 0.26019 | 0.22652 | 0.21129 | 0.1945 | 0.20035 | 0.20414 | 0.21607 | 0.23162
Time 6.7 h | -0.010233 | -0.013175 | -0.028274 | -0.022743 | -0.0054069 | -0.0034808 | -0.00042223 | 0.0044144 | 0.01012 | 0.014066 | 0.018998 | 0.024035 | 0.030298 | 0.037789 | 0.047292 | 0.058826 | 0.073223 | 0.090722 | 0.11222 | 0.13839 | 0.17098 | 0.21029 | 0.25838 | 0.3172 | 0.3871 | 0.46963 | 0.56416 | 0.66707 | 0.7712 | 0.86582 | 0.93655 | 0.97094 | 0.96234 | 0.91434 | 0.83635 | 0.74343 | 0.64608 | 0.55833 | 0.47266 | 0.39948 | 0.33451 | 0.29245 | 0.25353 | 0.23378 | 0.20871 | 0.19984 | 0.1997 | 0.2015 | 0.21198 | 0.22733 | 0.24895
Time 6.8 h | -0.0099355 | -0.010961 | -0.025677 | -0.02165 | -0.0043092 | -0.0020771 | 0.001197 | 0.0065657 | 0.012802 | 0.017476 | 0.023499 | 0.029632 | 0.037339 | 0.046572 | 0.058415 | 0.072736 | 0.09063 | 0.11238 | 0.13935 | 0.17222 | 0.21322 | 0.26261 | 0.32283 | 0.39563 | 0.48042 | 0.57735 | 0.68277 | 0.7887 | 0.88235 | 0.9508 | 0.98058 | 0.96664 | 0.91424 | 0.83246 | 0.73667 | 0.63856 | 0.5471 | 0.46601 | 0.38946 | 0.33579 | 0.28239 | 0.25179 | 0.22417 | 0.21528 | 0.19822 | 0.19607 | 0.20091 | 0.21064 | 0.22634 | 0.24719 | 0.27274
Time 6.9 h | -0.0089346 | -0.010265 | -0.024797 | -0.020617 | -0.0027907 | -0.00025292 | 0.0035689 | 0.0097638 | 0.016718 | 0.022291 | 0.029711 | 0.037389 | 0.047034 | 0.05871 | 0.073704 | 0.091917 | 0.11467 | 0.14237 | 0.17678 | 0.21881 | 0.27107 | 0.33354 | 0.40861 | 0.49749 | 0.59727 | 0.70497 | 0.81179 | 0.90383 | 0.96632 | 0.98864 | 0.96657 | 0.90704 | 0.82145 | 0.72363 | 0.62477 | 0.53332 | 0.45336 | 0.38378 | 0.32751 | 0.28107 | 0.24243 | 0.21966 | 0.20661 | 0.19414 | 0.1959 | 0.20098 | 0.21192 | 0.22772 | 0.24915 | 0.2732 | 0.30165
Time 7 h | -0.0079524 | -0.0096422 | -0.024013 | -0.019274 | -0.00065527 | 0.0023502 | 0.0069457 | 0.014294 | 0.022332 | 0.029183 | 0.0386 | 0.048527 | 0.061028 | 0.076199 | 0.095709 | 0.11952 | 0.14917 | 0.18538 | 0.23021 | 0.28467 | 0.35158 | 0.43028 | 0.52255 | 0.62693 | 0.73646 | 0.8422 | 0.92939 | 0.98297 | 0.99229 | 0.95868 | 0.8898 | 0.79968 | 0.70081 | 0.60339 | 0.51292 | 0.43557 | 0.37221 | 0.3158 | 0.27694 | 0.24492 | 0.21667 | 0.20627 | 0.19841 | 0.20582 | 0.20492 | 0.21652 | 0.23273 | 0.2534 | 0.27813 | 0.30477 | 0.33401
Time 7.1 h | -0.0072225 | -0.0078733 | -0.021826 | -0.017104 | 0.0022508 | 0.0058654 | 0.011358 | 0.020219 | 0.029818 | 0.038452 | 0.050584 | 0.063481 | 0.079875 | 0.09973 | 0.12537 | 0.15659 | 0.19542 | 0.24264 | 0.30066 | 0.37027 | 0.45418 | 0.54985 | 0.65682 | 0.7684 | 0.8716 | 0.95211 | 0.9947 | 0.99196 | 0.94565 | 0.86934 | 0.77503 | 0.67549 | 0.5806 | 0.49486 | 0.41965 | 0.35895 | 0.30957 | 0.26685 | 0.23671 | 0.2202 | 0.21086 | 0.19933 | 0.20391 | 0.21158 | 0.22344 | 0.23974 | 0.26069 | 0.28443 | 0.31084 | 0.33736 | 0.36523
Time 7.2 h | -0.0058046 | -0.0061413 | -0.019458 | -0.014478 | 0.0061524 | 0.010622 | 0.01737 | 0.028317 | 0.039983 | 0.051176 | 0.06699 | 0.083978 | 0.10566 | 0.13187 | 0.16569 | 0.20672 | 0.25752 | 0.3187 | 0.39276 | 0.47969 | 0.58074 | 0.68993 | 0.80068 | 0.9005 | 0.97199 | 1.0024 | 0.98623 | 0.92908 | 0.84405 | 0.74734 | 0.64904 | 0.55798 | 0.47566 | 0.40517 | 0.34692 | 0.30148 | 0.26468 | 0.23973 | 0.21784 | 0.21142 | 0.20723 | 0.20722 | 0.21754 | 0.22399 | 0.24796 | 0.26887 | 0.29214 | 0.31742 | 0.3425 | 0.36748 | 0.39208
-25.5 °C
Top electrode
Chart: Temperature /oC
| -40 | -39.1 | -38.299999999999997 | -37.4 | -36.6 | -35.5 | -34.799999999999997 | -33.9 | -33.1 | -32 | -31 | -30.2 | -29 | -28 | -27 | -26 | -25 | -24.1 | -23.1 | -22.2 | -21.1 | -20.2 | -19.3 | -18.5 | -17.899999999999999 | -17.100000000000001 | -16.600000000000001 | -16 | -15.5 | -15.2 | -14.7 | -14.5
Re-heating | 0.94718 | 0.95071 | 0.95411 | 0.95727 | 0.96001 | 0.96248 | 0.96474 | 0.96681 | 0.9687 | 0.97069 | 0.97255 | 0.97418 | 0.97558 | 0.97677 | 0.97768 | 0.97839 | 0.97899 | 0.97945 | 0.97971 | 0.97976 | 0.97957 | 0.97914 | 0.97846 | 0.97766 | 0.97681 | 0.97596 | 0.97513 | 0.97408 | 0.97303 | 0.97213 | 0.97127 | 0.97048
-1.00 x 10-4 | -5.20 x 10-3 | 9.19 x 10-1
-1.00 x 10-4 | -5.20 x 10-3 | 9.19 x 10-1
Polynomial coefficient from calibration
-14.5 °C
-40 °C
Top electrode
Temperature calibration of
III

## Slide 23: Objective
Compensation of during primary drying
IV

## Slide 24: Objective
Chart: Time / h
During primary drying, is attributed to both the loss of ice and product temperature; therefore, it requires a standardization factor () for temperature compensation:
The expression for can be re-written in terms of the polynomial coefficients (slide 22):
Values of during primary drying are then standardized to the reference temperature by dividing by to give a standardized peak amplitude of
and are the peak amplitudes at temperatures () and reference temperature () during the re-heating ramp. In this presentation, a temperature of -20 °C is used as the reference temperature value
The standardized is defined as
Chart: Time / h
Compensation of during primary drying
IV

## Slide 25: Objective
Calibration of for ice layer height
V
Filling water into TVIS vial
Freezing the sample
In-line TVIS measurement
Calibration plot ( vs temperature)
Calibration plot (𝐶_𝑃𝐸𝐴𝐾^″ vs temperature)
Thawing the sample
Identifying using simple peak finding
Identifying 𝐶_𝑃𝐸𝐴𝐾^″ using simple peak finding
Filling water into TVIS vial
Freezing the sample
In-line TVIS measurement
Thawing the sample

## Slide 26: Objective
Calibration of for ice layer height
V
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
0 mm | 0.0013746 | 0.00016925 | -0.010913 | -0.01186 | -0.0023631 | -0.0014374 | -0.0010379 | 7.0431e-05 | 0.00087291 | 0.00194 | 0.0032132 | 0.0043643 | 0.00576 | 0.0072677 | 0.0090276 | 0.010941 | 0.013034 | 0.015075 | 0.017002 | 0.018574 | 0.019679 | 0.020215 | 0.020242 | 0.019765 | 0.018906 | 0.017718 | 0.016447 | 0.015119 | 0.013913 | 0.012933 | 0.011994 | 0.011229 | 0.01049 | 0.0096809 | 0.0090548 | 0.0087375 | 0.00867 | 0.0083255 | 0.008589 | 0.0084126 | 0.0081989 | 0.0089905 | 0.0096046 | 0.0087964 | 0.011076 | 0.012465 | 0.013829 | 0.015592 | 0.017693 | 0.020004 | 0.022882
1 mm | 0.0027725 | 0.0018515 | -0.0092505 | -0.0099074 | 0.00073702 | 0.0022699 | 0.0034919 | 0.004273 | 0.0077695 | 0.0099906 | 0.01276 | 0.015324 | 0.018246 | 0.021403 | 0.025051 | 0.028978 | 0.033133 | 0.037091 | 0.04047 | 0.04265 | 0.043373 | 0.042565 | 0.040531 | 0.037669 | 0.034467 | 0.031215 | 0.028188 | 0.025539 | 0.023299 | 0.021591 | 0.020147 | 0.019038 | 0.018056 | 0.01722 | 0.016663 | 0.01663 | 0.016575 | 0.016995 | 0.017358 | 0.018559 | 0.019496 | 0.021177 | 0.022859 | 0.027278 | 0.028264 | 0.03148 | 0.03562 | 0.040533 | 0.046256 | 0.05258 | 0.060127
2 mm | 0.0025556 | 0.0042062 | -0.0069189 | -0.0070531 | 0.0050612 | 0.0074392 | 0.0097845 | 0.012275 | 0.017396 | 0.021327 | 0.026469 | 0.031628 | 0.037888 | 0.045042 | 0.053699 | 0.063205 | 0.073305 | 0.082842 | 0.090906 | 0.096188 | 0.098255 | 0.096899 | 0.0928 | 0.086718 | 0.079692 | 0.07254 | 0.065856 | 0.059996 | 0.055189 | 0.051295 | 0.048301 | 0.045995 | 0.04407 | 0.042496 | 0.041423 | 0.04147 | 0.042299 | 0.043522 | 0.045149 | 0.048416 | 0.051491 | 0.056861 | 0.062123 | 0.069787 | 0.075482 | 0.083507 | 0.092027 | 0.1012 | 0.11189 | 0.12216 | 0.13382
4 mm | 0.0098217 | 0.012857 | 0.0029162 | 0.0044657 | 0.021923 | 0.027475 | 0.03396 | 0.04256 | 0.053595 | 0.063826 | 0.077317 | 0.090942 | 0.1071 | 0.12594 | 0.14869 | 0.17399 | 0.20175 | 0.22964 | 0.25581 | 0.27689 | 0.29031 | 0.29404 | 0.28863 | 0.27556 | 0.25765 | 0.23771 | 0.21783 | 0.19948 | 0.18328 | 0.16931 | 0.15713 | 0.14631 | 0.13552 | 0.12507 | 0.11537 | 0.10625 | 0.097516 | 0.09077 | 0.085436 | 0.080247 | 0.074581 | 0.074997 | 0.072017 | 0.074699 | 0.073925 | 0.076692 | 0.080183 | 0.084679 | 0.090239 | 0.0972 | 0.1056
5 mm | 0.012924 | 0.016644 | 0.0043205 | 0.0094563 | 0.029609 | 0.036832 | 0.043438 | 0.058294 | 0.073283 | 0.087891 | 0.10754 | 0.12734 | 0.15165 | 0.17954 | 0.21314 | 0.25045 | 0.29123 | 0.33235 | 0.37121 | 0.40313 | 0.42452 | 0.43239 | 0.42664 | 0.40904 | 0.38343 | 0.35365 | 0.3231 | 0.29438 | 0.26853 | 0.24642 | 0.22701 | 0.20991 | 0.19402 | 0.17831 | 0.16442 | 0.15249 | 0.14204 | 0.13185 | 0.12652 | 0.1191 | 0.11251 | 0.11106 | 0.10999 | 0.10922 | 0.11194 | 0.11425 | 0.11948 | 0.12542 | 0.13331 | 0.14182 | 0.15289
6 mm | 0.016613 | 0.020933 | 0.0091737 | 0.015165 | 0.038068 | 0.04686 | 0.055308 | 0.072766 | 0.089943 | 0.10644 | 0.12816 | 0.14963 | 0.17566 | 0.20549 | 0.24169 | 0.28233 | 0.32749 | 0.37416 | 0.41944 | 0.45806 | 0.48528 | 0.49695 | 0.49191 | 0.47172 | 0.44084 | 0.40403 | 0.36577 | 0.32961 | 0.2979 | 0.27153 | 0.24965 | 0.23216 | 0.21743 | 0.20358 | 0.19134 | 0.18108 | 0.16982 | 0.15981 | 0.15095 | 0.14309 | 0.13722 | 0.12983 | 0.12612 | 0.12392 | 0.12523 | 0.12723 | 0.13111 | 0.13663 | 0.14387 | 0.15196 | 0.16262
7 mm | 0.019664 | 0.024611 | 0.012967 | 0.019785 | 0.045059 | 0.055353 | 0.065722 | 0.086411 | 0.1067 | 0.1266 | 0.15232 | 0.17875 | 0.21031 | 0.24651 | 0.29062 | 0.34037 | 0.39602 | 0.45396 | 0.51107 | 0.56056 | 0.59714 | 0.61478 | 0.61172 | 0.58909 | 0.55206 | 0.50663 | 0.45873 | 0.41313 | 0.37261 | 0.33899 | 0.31092 | 0.28831 | 0.26853 | 0.25064 | 0.23364 | 0.21847 | 0.20334 | 0.18982 | 0.17493 | 0.16472 | 0.15562 | 0.1442 | 0.14214 | 0.13187 | 0.13682 | 0.13729 | 0.14081 | 0.14603 | 0.15301 | 0.16257 | 0.17436
8 mm | 0.022191 | 0.024302 | 0.015888 | 0.023391 | 0.050987 | 0.062633 | 0.0751 | 0.099163 | 0.1228 | 0.14637 | 0.17727 | 0.20917 | 0.24729 | 0.29108 | 0.34437 | 0.40439 | 0.47128 | 0.5404 | 0.60767 | 0.66462 | 0.70504 | 0.72217 | 0.71465 | 0.68435 | 0.63803 | 0.58285 | 0.52571 | 0.47208 | 0.42452 | 0.38537 | 0.35336 | 0.32731 | 0.30445 | 0.28392 | 0.26444 | 0.24623 | 0.23159 | 0.21324 | 0.198 | 0.18347 | 0.16657 | 0.16468 | 0.15801 | 0.16092 | 0.15041 | 0.15193 | 0.15353 | 0.1576 | 0.16308 | 0.17249 | 0.18304
9 mm | 0.025178 | 0.02843 | 0.020221 | 0.028243 | 0.058241 | 0.071385 | 0.085782 | 0.11285 | 0.13953 | 0.1659 | 0.20185 | 0.2382 | 0.28153 | 0.331 | 0.3909 | 0.45815 | 0.53273 | 0.60911 | 0.68318 | 0.74564 | 0.78956 | 0.80825 | 0.80044 | 0.76808 | 0.71858 | 0.65905 | 0.59628 | 0.53734 | 0.48422 | 0.44063 | 0.40484 | 0.37572 | 0.35241 | 0.33355 | 0.316 | 0.3018 | 0.287 | 0.27231 | 0.25696 | 0.23941 | 0.22479 | 0.21074 | 0.20058 | 0.19055 | 0.18084 | 0.17458 | 0.17198 | 0.17099 | 0.17232 | 0.17616 | 0.18292
10 mm | 0.02931 | 0.032954 | 0.025296 | 0.034295 | 0.06725 | 0.078945 | 0.09936 | 0.13039 | 0.16103 | 0.19132 | 0.23267 | 0.27445 | 0.32415 | 0.38106 | 0.44994 | 0.52754 | 0.61349 | 0.70185 | 0.78726 | 0.85846 | 0.90721 | 0.92497 | 0.91033 | 0.86657 | 0.80291 | 0.72859 | 0.65233 | 0.5809 | 0.51827 | 0.46644 | 0.42508 | 0.39319 | 0.36784 | 0.34779 | 0.33105 | 0.31867 | 0.30683 | 0.29229 | 0.27976 | 0.26514 | 0.24949 | 0.23653 | 0.23008 | 0.23378 | 0.21063 | 0.2058 | 0.20258 | 0.20255 | 0.20419 | 0.20957 | 0.21751
11 mm | 0.030885 | 0.033883 | 0.025794 | 0.035193 | 0.069015 | 0.081061 | 0.10236 | 0.13477 | 0.16686 | 0.19878 | 0.24247 | 0.28687 | 0.3397 | 0.40029 | 0.47401 | 0.55745 | 0.65108 | 0.74884 | 0.84577 | 0.92979 | 0.99174 | 1.0211 | 1.015 | 0.97487 | 0.9107 | 0.83062 | 0.74636 | 0.66785 | 0.59703 | 0.53979 | 0.48998 | 0.45281 | 0.42298 | 0.39929 | 0.37893 | 0.36513 | 0.35333 | 0.34194 | 0.3318 | 0.31863 | 0.3062 | 0.29458 | 0.28223 | 0.28972 | 0.26381 | 0.25593 | 0.24876 | 0.2448 | 0.24301 | 0.2428 | 0.24608
12 mm | 0.032482 | 0.036695 | 0.028707 | 0.038445 | 0.073987 | 0.087194 | 0.10992 | 0.14456 | 0.17825 | 0.21335 | 0.26054 | 0.30844 | 0.36552 | 0.43109 | 0.51129 | 0.60208 | 0.70471 | 0.81264 | 0.92018 | 1.0145 | 1.0847 | 1.1186 | 1.1125 | 1.0687 | 0.99632 | 0.90626 | 0.81153 | 0.72126 | 0.64 | 0.57268 | 0.51614 | 0.47171 | 0.43528 | 0.40514 | 0.38388 | 0.36905 | 0.35669 | 0.3488 | 0.34184 | 0.33695 | 0.32609 | 0.32647 | 0.31616 | 0.32781 | 0.30771 | 0.30488 | 0.29988 | 0.29784 | 0.29674 | 0.29804 | 0.30238
Chart: -C″PEAK / pF
| 2.0282321658632974E-2 | 4.3255845536608106E-2 | 9.8028305551333794E-2 | #N/A | 0.29349694651821778 | 0.43151334859090174 | 0.49610799625102509 | 0.61477274370953339 | 0.72094798390835635 | 0.80705883860768424 | 0.92264736847743656 | 1.0206764847121299 | 1.118562503327631 | 1.1888173213507329 | 1.3470508093994198 | 1.4443671859293801 | #N/A | #N/A
 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | #N/A | #N/A
Trendline | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
0 mm
15 mm

## Slide 27: Objective
Estimation of ice layer height during primary drying
VI

## Slide 28: Objective
At 2.4 h into primary drying
= 0.863 pF
Ice height = 9.7485 x 0.863 + 1.0462
= 9.459 mm (from the bottom edge of TE)
Ice front height = 9.459+(2+5+3)
= 19.46 mm
Bottom electrode (BE)
5 x 19 mm
3 mm gap
Freeze Dryer Shelf
Chart: -C″PEAK / pF
| 2.0282321658632974E-2 | 4.3255845536608106E-2 | 9.8028305551333794E-2 | #N/A | 0.29349694651821778 | 0.43151334859090174 | 0.49610799625102509 | 0.61477274370953339 | 0.72094798390835635 | 0.80705883860768424 | 0.92264736847743656 | 1.0206764847121299 | 1.118562503327631 | 1.1888173213507329 | 1.3470508093994198 | 1.4443671859293801 | #N/A | #N/A
 | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | #N/A | #N/A
Trendline | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
3 mm gap
Top electrode (TE)
15 X 19 mm
At -20 °C
Freeze drying shelf
2 mm
5 mm
3 mm
0.863 pF
19.46 mm
10 mm
9.46 mm
Gradient of the line ()
Estimation of ice layer height during primary drying
VI

## Slide 29: Objective
Surrogate drying rate can be estimated in terms of decreasing ice height
The dependency of on the ice cylinder height in linear region
Chart: Time / h
20.50 mm
Estimation of ice layer height during primary drying
VI

## Slide 30: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII

## Slide 31: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
-32.3 ˚C
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
2 mm
5 mm
3 mm
9.46 mm
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C

## Slide 32: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
0 mm
-30.6 ˚C
-32.3 ˚C

## Slide 33: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C

## Slide 34: Objective
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII
Freeze drying shelf
Chart: Temperature (𝑇) / oC
| -33.119466634243032 | -32.321208025650492 | -30.594741732648018 | -29.835299081767161
TE_BE_predict | 19.46 | 14.73 | 4.5 | 0
Chart: Time / h
4.5 mm
-32.3 ˚C
14.73 mm
Chart: Time / h
Ice height 19.46 mm at 2.4 h
-30.6 ˚C
At 2.4 h
Ice height for = () = 14.73 mm
Ice height for = () = 4.50 mm
2 mm
5 mm
3 mm
9.46 mm
-33.1 ˚C
-29.8 ˚C
0 mm
19.46 mm
-30.6 ˚C
-32.3 ˚C
Ice Temperature
At interface (, 19.46 mm) = = -33.1 ˚C
At vial’s base (, 0 mm) = = -29.8 ˚C

## Slide 35: Objective
The product temperature at ice interface predicted by using a 2-points temperature extrapolation close to the temperature of ice vapour at chamber pressure of 270 µbar ()
Chart: Time / h
Shelf Temperature
Chart: Time / h
Temp. constant
(= -33.1± 0.05˚C )
Temp. constant
(= -33.1± 0.05˚C )
Prediction ice temperatures at (i) sublimation interface () and (ii) vial’s base () including qualification TVIS technique ()
VII

## Slide 36: Objective
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII

## Slide 37: Objective
Drying rate is based on the assumption of a planar sublimation front
The change in ice cylinder height () can be equated to the change in ice volume ()
Where is internal radius of vial and is internal cross section area of vial ()
Ice volume can be converted to ice mass () by multiplying with ice density ()
Hence; drying rate () can be expressed by
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII

## Slide 38: Objective
An average surrogate drying rate calculation
  Ice density at -32 ˚C = 0.920 gcm-3
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 0 h () = 20.50 mm
  Ice height at 4.9 h ( = 16.68 mm
 | 0.27 g/h
 | 0.25 g/h
 | 0.27 g/h
 | 0.25 g/h
Chart: Time / h
20.50 mm
Comparison of TVIS drying rate () with gravimetric method (weight loss)
VIII

## Slide 39: Objective
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period
IX

## Slide 40: Objective
Chart: Time / h
Chart: Time / h
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 41: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
Drying rate during the steady state
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 42: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 43: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 44: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 45: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 46: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
  Ice height at 2.8 h ( = 18.98 mm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 47: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
~31.5 ˚C
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
  Ice height at 2.8 h ( = 18.98 mm
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 48: Objective
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
19.94 mm
Drying rate during the steady state
  Ice density at -32˚C = 0.920 gcm-3
  (Calculated ice temperature between
  Internal vial diameter (VC010-20C) = 2.21 cm Cross-section area () = 3.80 cm2
  Ice height at 2 h () = 19.94 mm
  Ice height at 2.8 h ( = 18.98 mm
~32 ˚C
TVIS parameters used for determination:
= 0.42 g·h-1
= -29.8˚C
Determination (i) the drying rate () and (ii) ice base temperature () during the steady state period for heat transfer coefficient ( calculation
IX

## Slide 49: Objective
Heat transfer coefficient ( calculation
X

## Slide 50: Objective
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 51: Objective
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
is the latent heat of sublimation of ice (2844 Jg-1 or 679.7 cal g-1) and is external cross-sectional area of the base of the TVIS vial (4.62 cm2)
Pikal, et al. (1984)
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 52: Objective
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
is the latent heat of sublimation of ice (2844 Jg-1 or 679.7 cal g-1) and is external cross-sectional area of the base of the TVIS vial (4.62 cm2)
Pikal, et al. (1984)
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 53: Objective
Chart: Pressure (mTorr)
| 203.25
TVIS | 0.00057283
Tchessalov S (2017) | 0.00075778 | 0.00057449 | 0.00057449 | 0.0002011 | 0.00030003 | 0.00038954 | 0.0004709 | 0.00054519 | 0.00061327 | 0.00016575 | 0.0002352 | 0.00036064 | 0.00056849
TVIS @ 270 bar)
5.73 x 10-4
is the latent heat of sublimation of ice (2844 Jg-1 or 679.7 cal g-1) and is external cross-sectional area of the base of the TVIS vial (4.62 cm2)
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
Shelf Temperature, (K) | 273.3
Vial’s base Temperature, (K) | 243.3
Parameters | TVIS
Drying rate at steady state (g/h) (2-2.8 h into primary drying) | 0.42
 | 273.3
 | 243.3
Heat transfer coefficient ( calculation
X
Pikal, et al. (1984)
Notes: Freezing -40 C
Annealing -10 C
Drying -20 C

## Slide 54: Additional comments
Qualification of steady state heat transfer mechanisms

## Slide 55: A single vial technique Pikal, et al. (1984)
The mean sublimation rate was calculated from the mass of ice sublimed and the time required for sublimation.
Notes: [1] M. Pikal, M. Roy, S. Shah, Mass and heat transfer in vial freeze‐drying of pharmaceuticals: Role of the vial, J. Pharm. Sci. 73 (1984) 1224-1237.
[2] S. Hibler, C. Wagner, H. Gieseler, Vial freeze-drying, part 1: New insights into heat transfer characteristics of tubing and molded vials, J. Pharm. Sci. 101 (2012) 1189-1201.
[3] B. Scutella, S. Passot, E. Bourles, F. Fonseca, I.C. Trelea, How Vial Geometry Variability Influences Heat Transfer and Product Temperature During Freeze-Drying, J. Pharm. Sci. 106 (2017) 770-778.

## Slide 56: Assumption for determination
How do we know that the heat transfer mechanisms are constant up to 25% loss of ice mass?
If the heat transfer mechanisms change because of ice- glass interface contact or because of the change of ice shape (surface area) then surely heat transfer coefficient will change?
It requires a technique to qualify when the heat transfer mechanisms change
So can TVIS demonstrate when ice leaves the glass wall interface?

## Slide 57: Limitation of TVIS System ?
Chart: Time / h
Chart: Time / h
Temp. constant = -33.1± 0.05˚C
= -29.8± 0.03˚C
~32 ˚C
I
II

## Slide 58: Discussion
Decrease in suggests that the temperature may be decreasing after the steady state period, contrary to accepted knowledge that the temperature starts to increase owing to a reduction in drying rate and hence the degree of self cooling
Decrease in is more likely to be due to a change in the ice-glass contact associated with a change in the shape of the ice cylinder.
Conclusion
The period for determining the drying rate should be decreased from 25% ice loss to 15% for TVIS to give reliable estimates for Kv
Opportunity to cycle through shelf temperature and chamber pressure to create the design space for Kv determinations as a function of shelf position.

## Slide 59: Limitations
and parameters rely on intimate contact of ice cylinder with glass wall
(100 kHz) parameter does not dependent on contact and can be used for end point but relationship between (100 kHz) ice constant is non-linear
Cable length limited to 1m at present
C-TVIS not compatible with front loading system
Incompatible with TCs in same TVIS vial (use fibre optic sensors – INFAP)

## Slide 60: Future Work
Development dryer mapping of sublimation characteristics
  heat transfer coefficients ()
  dry layer resistance ()
Instrument Development
  Contact C-TVIS instrument (2018)
  Non-contact TVIS (2018-19)
    Micro-well screening
    Vial clusters in batch FD
  TVIS - Shuttle (2019-20)

## Slide 61: Acknowledgements, Recent Projects & Collaborators
De Montfort University, School of Pharmacy
  Evgeny Polygalov: co-inventor of TVIS instrument
  Yowwares Jeeraruangrattana. PhD student
  Bhaskar Pandya. PhD student
  Irina Ermolina. Senior Lecturer
Biopharmaceutical Stability at
Room Temperature
Analytical Technologies for the
Stabilization of Biopharmaceuticals
Government Support for industry
LyoDEA
Lyophilization process analytics
By dielectric analysis

## Slide 62: Thank you